## RT-PCR-based Molecular Tests
Table SI. RT-PCR-based molecular tests currently available on the market and their main technical features.
Kit Name | Type | Target gene | Sensitivity | Specificity | LoD | Company/Supplier | Website | Time
AcculaTM SARS-CoV-2 Test | Nucleic acid amplification test (NAAT) | N | N/A | N/A | 150 copies/ml | Mesa Biotech, Inc. | https://www.mesabiotech.com | 30 min
Xpert Xpress SARS-CoV-2 | RT-qPCR-based method | N2 - E | 97.8% | 95.6% | 0.0200 PFU/ml | Cepheid | https://www.cepheid.com | 30-45 min
Amplidiag® COVID-19 | RT-qPCR-based method | N - orf1ab | 98.9% | 98.8% | N/A | Mobidiag | https://mobidiag.com | <3 h
QIAstat-Dx Respiratory Panel 2019-nCoV | RT-qPCR-based method | E - RdRp | N/A | N/A | 400 copies/ml | Qiagen | https://www.qiagen.com | 1 h
NeuMoDx SARS-CoV-2 Assay | RT-qPCR-based method | Nsp2 - N | 100% | 100% | 150 copies/m | NeuModDx molecular | https://www.neumodx.com | N/A
Abbott RealTime SARS-CoV-2 | RT-qPCR-based method | RdRp - N | 100% | 100% | 100 copies/ml | Abbott Laboratories | https://www.molecular.abbott | ~4 h
Cobas® SARS-CoV-2 | RT-qPCR-based method | ORF1 a/b - E | 100% | 100% | 0.009 TCID50/ml | Roche | https://diagnostics.roche.com | 3 h
Vivalytic CoV Viral SARS-CoV-2 test | RT-qPCR-based method | ORF1 a/b - E | 100% | 100% | N/A | Bosch - Vivalytic | https://www.bosch-vivalytic.com | 38 min
UltraGene Combo2Screen SARS-CoV-2 Assay | RT-qPCR-based method | E - N | 100% | 100% | 1x10-6 TCID50/ml | ABL SA Group | https://www.ablsa.com | 2 h
RealStar® SARS-CoV-2 RT-PCR Kits | RT-qPCR-based method | E - S | N/A | N/A | 10 copies/PCR | Altona Diagnostics | https://altona-diagnostics.com | 1 h
Bosphore Novel Coronavirus (2019-nCoV) Detection Kit | RT-qPCR-based method | ORF1 a/b - E | N/A | 100% | <13 copies/μl | Anatolia Diagnostics | http://www.anatoliageneworks.com | 1.5 h
iAMP® COVID-19 Detection Kit | Real-time fluorescent reverse transcription isothermal assay | ORF1 a/b - N | 100% | 100% | 10 copies/μl | Atila BioSystems | https://atilabiosystems.com | <1.5 h
BioGX SARS-CoV-2 Reagents (for BD MAX™ System) | RT-qPCR-based method | N1 - N2 | 100% | 100% | 40 GCE/ml | Becton, Dickinson | https://www.bd.com | 3 h
Real-Time Fluorescent RT-PCR Kit for Detecting SARS-CoV-2 | RT-qPCR-based method | ORF1 a/b | 88.1 | 99.6% | 150 copies/ml | BGI Genomics | https://www.bgi.com | 4 h
EurobioPlex SARS-CoV-2 Multiplex | RT-qPCR-based method | RdRp - N | >99% | 98.6% | N/A | Eurobio Scientific | https://www.eurobio-scientific.com | 1.15 h
EURORealTime SARS-CoV-2 | RT-qPCR-based method | ORF1 a/b - N | 100% | 98.2% | 1 copy/μl | EUROIMMUN; PerkinElmer, Inc. | https://www.euroimmun.com | ~1.5 h
Panther Fusion® SARS-CoV-2 Assay | RT-qPCR-based method | ORF1 a/b | 100% | 100% | 1x10-2 TCID50/ml | Hologic, Inc. | https://www.hologic.com | 2-4 h
TaqPath COVID-19 CE-IVD RT-PCR Kit | RT-qPCR-based method | ORF1 a/b - N - S | 100% | 100% | 10 GCE/reaction | Applied Biosystems; Thermo Fisher Scientific, Inc. | https://www.thermofisher.com/it/en/home.html | <2 h
OPTI® SARS-CoV-2 RT PCR Test | RT-qPCR-based method | N1 - N2 | 100% | 100% | 1 copy/μl | OPTI Medical Systems, Inc. | https://www.optimedical.com | 2-3 h
COVID-19 genesig® Real-Time PCR assay | RT-qPCR-based method | RdRp | 94.7% | 100% | 0.33 copies/μl | Primerdesign Ltd. | https://www.genesig.com | 2-4 h
Lyra® SARS-CoV-2 Assay | RT-qPCR-based method | Pp1ab | 100% | 100% | 0.8 copies/μl | Quidel Corporation Headquarters | https://www.quidel.com | 75 min
Novel Coronavirus (2019-nCoV) Nucleic Acid Diagnostic Kit (PCR-Fluorescence Probing) | RT-qPCR-based method | ORF1 a/b - N | 94.3% | 99% | 200 copies/ml | Sansure Biotech, Inc. | http://eng.sansure.com.cn | ~1.5 h
STANDARD M nCoV Real-Time Detection kit | RT-qPCR-based method | ORF1ab - E | 100% | 100% | 0.5 copies/μl | SD Biosensor Inc | http://sdbiosensor.com/ | ~1.5 h
AllplexTM2019-nCoV Assay | RT-qPCR-based method | RdRp - N - E | 100% | 93.07% | 4.167 copies/ml | Seegene, Inc | http://www.seegene.com/ | <2 h
2019-Novel Coronavirus (2019-nCoV) RT-PCR Detection Kit | RT-qPCR-based method | ORF1ab - N - E | 99.53% | 97.92% | 300 copies/ml | Shanghai Fosun Long March Medical Science Co., Ltd. | http://en.lm-diagnostics.com.cn | ~2 h
YouSeq Multiplex Covid19 qPCR Kit | RT-qPCR-based method | RdRp - N - E | 100% | 100% | 0.7 copies/μl | YouSeq Ltd. | https://youseq.com | 2-4 h
VitaPCR SARS-CoV-2 assay | RT-qPCR-based method | N | 100% | 100% | 2.73 copies/μl | Credo Diagnostics Biomedical Pte Ltd. | www.credodxbiomed.com | ~20 min
LYRA SARS-CoV-2 assay | RT-qPCR-based method | pp1ab | 100% | 100% | 8.00E-01 copies/μl | Diagnostic Hybrids, Inc. Quidel Corporation | https://www.quidel.com | <2 h
Simplexa COVID-19 Direct | RT-qPCR-based method | ORF1ab - S | 100% | 100% | 500 copies/ml | DiaSorin Molecular LLC | https://molecular.diasorin.com | ~1.5 h
ARIES® SARS-CoV-2 Assay | RT-qPCR-based method | ORF1ab - N | 100% | 100% | 5.48 E-02 TCID50/ml | Luminex Corporation | https://www.luminexcorp.com | ~2 h
MiRXES FORTITUDE KIT 2.1 | RT-qPCR-based method | ORF1ab | 100% | 100% | 10 copies/reaction | MiRXES Pte Ltd. | https://mirxes.com | <2 h
E, envelope gene; Nsp2, non-structural protein 2 gene; N, nucleocapside gene; ORF1ab, open reading frame 1ab gene; pp1ab, polyprotein 1ab gene; Rdrp, RNA-dependent RNA polymerase gene; S, Spike gene; GCE, genome copy equivalent; LoD, limit of detection; N/A, not available; TCID50, median tissue culture infectious dose.